SHR Fast Hair Removal Machine

Operation manual

Chapter 1. INTRODUCTION

Front view

Back view

Chapter 2. Installation

1. In stall the pole and the handle holder.

- Connect handpiece

2.1.1 Press down the 2 buttons firstly.

There are two buttons on both above and below of each handpiece. The two buttons may be in the upspring or press-down state. Before connect it onto the machine, it must be in the press-down state. So press down the 2 buttons firstly.

2.1.2 Connect it in the right direction.

There are two parts of the connectors: machine part and handpiece part. Please look at the following pictures. A1 connects to B1, and A2 connects to B2. It cannot plug if connecting A1 to B2.

2.1.3 Use some power of your hands to plug it. When you plug one well, you can hear the click – the button springs and it locks. Then you can try to pull the handpiece connector slightly to check whether or not it locks well. Then hang up the two handpieces onto handpiece holders.

After installation:

- Fill water

- Open the overflow hole

The small white plastic screw shall never be plugged again on the machine. Please read the following sticker.

Pure water or distilled water is needed.

Back view

Press down the button on the Water-In hole.

Insert the water pipe with plug and add water via funnel.

Add water until water runs from Water Overflow hole.

Press the button on the Water-In mouth, the water pipe pops out automatically.

Chapter 3. Operation

1. Equipment operation (treatment process):

(1) The requirements for operator:

Apparatus for the operator should understand E light treatment of professional knowledge, the training of qualified medical professionals or beautician operate the machine.

(2) Operation steps:

① Add water

Firstly you should open water tank cover with funnel to add distilled water or pure water, prohibit the use of mineral water, and then boot the machine, water level will drop, then the intake, water flow circulation normal.

Boot the machine

(1) You need switched on the power supply

(2) In the emergency stop switch turned on, right turning key switch, that machine begin to work.

(3) That you can see the following screen interface:

① Turn on the screen and enter Figure one as below:

Figure one

Click SHR.

②Treatment interface

Figure two

Click Handle cool levels into full levels (from gray  to ) and the sapphire becomes cold in 1 or 2 seconds.

Pulse width: how long a shot work on skin. 1~7ms adjustable

Frequency: how many shots in one second. 1~10Hz adjustable

Energy:

- when the frequency is 1hz/2hz/3hz/4hz, the energy is 1-50J adjustable

- when frequency is 5hz, energy is 1-35J adjustable

- when frequency is 6hz, energy is 1-30J adjustable

- when frequency is 7hz, energy is 1-25J adjustable

- when frequency is 8hz, energy is 1-20J adjustable

- when frequency is 9hz, energy is 1-15J adjustable

- when frequency is 10hz, energy is 1-10J adjustable

Figure three

When to adjust pulse width or Frequency, you need click the figure into red color:

Figure four

Then click  to increase it or  to decrease it. Then click the red figure into white color again, or other settings are not able to change.

When you test the machine, you can use pulse width 3ms, frequency 3hz and energy 5J. Then raise the settings slowly according to actual situation.

When settings are ready, press "Ready” button to trigger the treatment head and to charge the capacitances automatically. Press the red button and it shots.

When treatment finishes, press "Standby” button to turn off the head.

Status prompt window

When the machine is Standby, it shows. When click Ready, it shows .

It may change between  and  during fast shots. It is normal.

When the water cycling is bad, the machine alarms and it shows .

When water temperature is too high, the machine alarms and it shows .

Chapter 4. Treatment

Before shooting, operator shall wear OD4+ protective glasses and patient shall wear eye patch.

The treatment area should be coated with a thin layer of coupling agent (such as cold gel).

Fix settings and trigger the handle, let sapphire touch skin tightly, keep pressing the handle switch and slide the handle on skin. The slide speed depends on how fast the frequency is. You need to ensure all the treatment area is covered by the light.

Release the handle switch before the handle leaves skin.

Attention after operation

1、Prevent from solarization: As the solarization may spur the rebirth of melanin, in 30 days after operation, patient should apply suntan oil with SPF 15 or sun block, and hold parasol or wear sunbonnet at outdoor place.

2、Face-painting: Face-making is not recommended after

treatment. A few days after operation, patients are not allowed apply chemical cosmetic or powdery makeup while biological skin-protection products are better. If symptoms like red turgescence or effusion, patient should report to doctors immediately and stop face-paiting. Take some Vitamin C and Vitamin E will help the treatment.

According to the normal physiological circle of human body, patient should go on with the next treatment 3 weeks after the operation. A period of treatment includes 3-5 treatments. After a period of treatment, typical flaws of skin like solar speckle, age pigment will be removed, and at the same time the skin will be tender, glossy and elastic.

Attention: Due to the constitutional difference, some clients may get vesicles, pigment diposition, red turgescence and gore. So patient could apply ice pack to the affected part for 30 minutes. Face- making is forbidden until 5 days later.

Contraindication

Cancer patients, hypotensives and diabetes sufferers, or people

With scar, cut or infected skin, abnormal diathesis or

Hyper susceptibility patients and people with irregular incretion are forbidden to undergo the operation.

Maintaining and Service

The equipment is constituted by precise optics instrument or electronic instrument. The maintaining and service of the machine should agree with the general regulation of precise optics instrument

Or electronic instrument. Avoid knocking, extrusion and vibration in placing and using it. Check the surface of lens after every operation, if there are splashing objects on it, dab them away with lens paper or cotton sticker with absolute alcohol

Skin-test (Essential)

Skin test is essential before the treatment for the safety operation. The detail is as following:

- Skin test should be performed before the treatment for the client when the first time. The test patch should be under the ears or other skins not influencing the facial beauty.

- Clean the test patch with pure water and wash- milk; apply a layer of cold coupling gel at the test patch. Take the treatment handle from the Supporter, and direct the light guide to the test patch. Trigger 2-3 shots (1/3 area of the spot size overlapping the treatment site is permitted.), and wait for 2-3 minutes to examine the test patch:

- If lightly swelling and/ or reddening at the test patch appears several minutes later, continue the treatment with the same settings of energy level and pulse type.

- If no visible changes at the patch appear, continue the test with 2j/cm² higher energy level. Re-perform the test for several times until the above situation happens.

- If the skin show adverse effect or excessive reddening and/or swelling at the test patch just after shot, continue the test with lower energy level.

- First, cut off the hair and leave only 1-2 mm. clean the test patch with pure water, and apply a layer of cold coupling gel at the test patch. Take the treatment handle from the supporter, and direct the light guide to the test patch. Trigger 2-3 shots (1/3 area of the spot size overlapping the treatment site is permitted.), and wait for 2-3 minutes to examine the test patch:

- If swelling and/or reddening around the hair follicles appear within a few minutes, continue the treatment with the same settings.

- If no visible changes appear around the follicles, continue the test with 2j/cm² higher energy level.

- If the skin shows adverse effect such as excessive reddening or swelling), lower the energy level.

Chapter 5. post-treatment care

No exposure to sunlight

Clients should use highest factor sunscreen and protect the treated area from exposure to sunlight, sunbed or other artificial tanning for one month following treatment. Tanning after treatment sessions may enhance melanin regeneration, which may result in hyper-pigmentation

Free from always touching and pressing

Clients must do not press and scratch the treated area against infection

Makeup immediately after the treatment is not allowed for better curative-effect, advising no use of makeup for four to five days following treatment. The skin is sensitive during this period. Remove the makeup when swelling and reddening appearing at the treated site against infection and pigmentation. In the mean time advise client take some

Food

Advising not eat the hot and spicy food because the skin is sensitive during this period.

Other

Advise that improve other causes of disease, such as adjusting the endocrine situation by taking Chinese herb of medicine, taking few medicines that include VitC and Vit E. According to the normal physiological process, clients should take the next treatment 21 days later. The whole treatment needs 3-5 times.

Chapter 6 ATTENTIONS OF OPENING THE SCREEN CASE

You may need to open the screen case for maintenance in future. It is better to have two persons to open it.

1. Unscrew the following 8 screws. At the same time, one person must hold the front part of the screen case; otherwise, the whole front part will drop down when all the screws are uninstalled.

Chapter 7 Technical Specifications

| SPECIFICATIONS FOR SHR HAIR REMOVAL MACHINE |
| --- |
| Type of machine | SHR |
| Wavelength | 640-950nm |
| Energy density | 1-50J/cm2 adjustable |
| Spot size | 10x40mm or 15x50mm |
| Pulse duration | 1-7ms |
| Pulse | Single |
| Shot frequency | 1-10HZ |
| Display | 10.4' Movable color LCD touch screen |
| Cooling system | Semiconductor+ air + water + cooling gel |
| Skin Cooling | -6°C ~ 0°C |
| Output power | 2500W |
| Voltage | 100~240V 50/60Hz |
| Multi software languages | Yes |
| Product size | 49x41x34cm(screen lays down) 49x41x54cm(screen stands up) |
| Package size | 65x53x54.5cm |
| Net weight | 26KG |
| Gross weight | 40KG |

# Text boxes

SHR handpiece

SHR handle connector

Key Switch

USB port

10.4’’ movable color touch screen

Quick Stop

Water Overflow hole

Water-In Hole and Water Discharge hole

Electrical power port

Machine part

A2

A1

Handpiece part

B2

B1

Water Overflow hole

Water-In Hole and Water Discharge hole

Electrical power port

Treatment mode

System setting

6 settings memories

Pulse width

Frequency

Counter

Status notification

Return Key

Ready /standby

Contact cooling

SHR energy adjustment

Beijing Renguang Weiye Science & Technology Development Co. Ltd